## Formulae
Common Constants & Variables | Symbol | Value | Units | Examples | Comments
Constant as per section 3.2.2 | Const322 | 838860800
Constant as per section 3.2.4.3 | Const3243 | 3763.74
Constant as per section 3.2.4.4 | Const3244 | 93.2068
Pulse constant on energy meter | MC | 3200 | imp/kWhr
Reference Voltage METERED | Un | 238.7 | Volts | 220.024 | As measured using multimeter (with reference load active)
Basic Current METERED | Ib | 7.14 | Ampere | 5.008 | As measured using multimeter (with reference load active)
Current Transformer Ratio | Nx | 2000 | Ratio | 0.001 | Primary:Secondary windings (set to 0.001 for 200uOhm shunt example)
Voltage circuit sampling voltage at Un | Vu | 270.9421 | mV |  | Calculated based on resistor network
Burden resistor | Rb | 11.6 | Ohm | 0.0002 | Dual 5.8Ohm burden resistors, what is correct value
#1 - Plconstant (see 3.2.2)
L line current circuit gain RegX2B | Gl | 1 |  | 24 | Based on lGain[2:0] translated to value 1/4/8/16/24
L line Current sampling voltage at Ib | Vl | 41.412 | mV | 1
PL Constant H+L (-> x21 & x22) |  | 1725806.1294
RegX2122 | PLconstLH | x001A556E |  |  | Check all variables above
#2 - Mmode Register (see 3.2.3)
#3 - Power Offset Compensation (see 3.2.4.1)
Done using software M90P {x}
#4 - L Line Gain (PF=1.0, current=Ib) (see 3.2.4.2) |  |  |  | 1b9e
L Line Current RMS Register | RegX48 | 7079 | HEX <====== | 1ba7 | Use x10d7 for Lgain & x1A58 for Igain
L Line Current RMS Measured | Cur_meaL | 28.793 | Ampere
Calibration Bench Error Ratio | _e | 303.3% |  |  | Difference/error Ib (multimeter) vs Irms (measured)
Lratio | Lratio | -0.75202
L line Gain Calibration Value |  | 40893.7083
RegX23 | Lgain | x9FBE |  |  | Check if Ib & RegX48 above is correct !!!!
#5 L Line Angle (PF=0.5L, F=50Hz, current=Ib) (see 3.2.4.3)
L-Line Phase Angle METERED | ?????? | 1.0095 | DEG <====== | 1.01 | Any suggestions on how to practically measure this????
L-Line Phase Angle Register | RegX4E | 000A | HEX <====== | 000A
L-Line Phase Angle Measured |  | 1 | Degrees |  | Converted to range -180 to +180
Calibration Bench Error Output | _el | 0.95%
Angle | Angle | -0.31426
L Line Angle Calibration Value |  | 35.7555
RegX24 | Lphi | 0024 |  |  | Check if Angle METERED & RegX4E is correct !!!
#6 - Startup/NoLoad, Power/Current Threshold (see 3.2.4.4)
| Ks | 0.004 | <========== | 0.004 | Where to we get/estimate this from ?????
Active Startup Power Threshold |  | 4183.2142
RegX27 | PStartTh | 1057 |  |  | Check value for Ks supplied above is correct
| Kn | 0.004 | <========== | 0.004 | Where to we get/estimate this from ?????
Active NoLoad Power Threshold |  | 4183.2142
RegX28 | PNolTh | 1057 |  |  | Check value for Kn supplied above is correct
| KsQ | 0.004 | <========== | 0.004 | Where to we get/estimate this from ?????
ReActive Startup Power Threshold |  | 4183.2142
RegX29 | QStartTh | 1057 |  |  | Check value for KsQ supplied above is correct
| KnQ | 0.004 | <========== | 0.004 | Where to we get/estimate this from ?????
ReActive NoLoad Power Threshold |  | 4183.2142
RegX2A | QNolTh | 1057 |  |  | Check value for KnQ supplied above is correct
#7.1 - L Line Voltage RMS Gain (see 3.2.4.5)
Ugain Default Value (hardcoded) | VlReg | 26400 |  |  | LEAVE, default
L Line Voltage RMS Gain Measured | Vol_mea(X49) | 236.86 | HEX <====== | 5c86
Voltage RMS Gain calibration |  | 26605.0832
RegX31 | Ugain | x67ED |  |  | Check if RegX31 above is correct !!!
#7.2 - L Line Current RMS Gain
L Line Current RMS Gain Default | Cur_defL | 31251
L Line Current RMS Gain Calculated |  | 7749.5273
RegX32 | Igain | 1E46 |  |  | Check if RegX48 above is correct !!!!
#8 - N Line Current RMS Gain (see 3.2.5.x)
N Line Current RMS Gain Default | Cur_defN | 30000 |  |  | LEAVE, default
N Line Current RMS Register | RegX68 | 30000 | HEX <====== | 7530
N Line Current RMS Measured | Cur_meaN | 196.608
N Line Current RMS Gain Calculated |  | 1089.4775
RegX25 | IgainN | 0441 |  |  | Check if RegX68 above is correct !!!
#9 - Current Offset compensation (see 3.2.4.1)
Done using software
Sequence of calibration:
Voltage
Current | Check power, should be right
PLconstHL
Note on F10: As per example 200u Ohm

## Sheet1
Item | Load (watt) | Type
Kettle |  | Resistive
Oil heater | 3000 | Resistive
incandescent lamp | 150 | resistive
Extraction fan | 45 | Resistive

## CT's
Manufacturer | Model | Sub-Model | Input | Output | Ratio | Type | Quantity
YHDC | SCT013 | 5A/1V | 0-5A | 0-1V |  | Voltage | 2
|  | 10A/1V | 0-10A | 0-1V |  | Voltage | 2
|  | 30A/1V | 0-30A | 0-1V |  | Voltage | 3
| ZHT103 |  | 0.3-5A |  | 1:1000 | Voltage | 5
PZCT |  |  | 0-50A | 0-10mA | 5000:1 | Current | 6